## Slide 1: Bi9393 Analytická cytometrie Lekce 3
Oddělení cytokinetiky
Biofyzikální ústav AV ČR, v.v.i.
Královopolská 135
612 65 Brno
e-mail: ksoucek@ibp.cz
tel.: 541 517 166
Karel Souček, Ph.D.
K. Souček Bi9393 Analytická cytometrie

## Slide 16: Particle Delivery: Hydrodynamic Focusing
Intensity
Count
Narrow particle focus = Narrow distribution
Laser Cross Sectional Area
Sample core is ‘pinched’ by fast flowing sheath
Sample volume ratios of 100 – 1000
Large ratios => low sample inputs
Resolution of particle populations
sheath
sheath
Hydrodynamic core
Conventional Instrumentation: Low Flow Rates (12µL/min)
Notes: To really understand what is different and unique about acoustic focusing, one first needs to understand the current most commonly used method for focusing, called hydrodynamic focusing.
In this method, the sample is injected into a stream of sheath fluid where it is completely surrounded by sheath, the sample core is then centered in the sheath fluid where the laser beam will then interact with the particles. Based on principles relating to laminar flow, the sample core remains separate but coaxial within the sheath fluid. The flow of sheath fluid accelerates the particles and restricts them to the center of the sample core. This process is known as hydrodynamic focusing. Where the sample core is narrow, the particles interact with the laser beam optimally and a narrow distribution results as seen in the histogram above.

## Slide 17: Particle Delivery: Hydrodynamic Focusing
Conventional Instrumentation: High Flow Rate (60µL/min)
Intensity
Count
Broad particle focus = Broad distribution
Increased sample input = increase core size
Particle distributions broadened, CVs increase
Instrument resolution decreased
Historically, low volumetric sample rates used (25 ml/min – 150 ml/min)
sheath
sheath
Hydrodynamic core
Laser Cross Sectional Area

## Slide 18
Attune® Acoustic Focusing Cytometer

## Slide 19: Acoustic Focusing = Better Precision
Acoustic focusing Module
sheath
sheath
sheath
sheath
Narrow particle focus = Narrow distribution
12 µL/min
1000 µL/min
Acoustic focusing of particles occurs prior to mixing with sheath fluid
Notes: Fundamentals of acoustic focusing in flow cytometry:
Picture in the middle is a comparison of a manifold taken from a traditional hydrodynamic focusing cytometer (left) compared to a manifold taken from the Attune Cytometer and represented “in action” by the cartoons directly to the left and right of the photograph. The picture on the far right shows a full acoustic device as well as the manifold at the top.
Here, regardless of the sample volume that is moved through the flow cell, cells are focused by the acoustic device prior to entering the manifold and, ultimately, the flow cell.
So, at low sample volumes (left) and high sample volumes (right), focus is maintained.
This also means that very little sheath is needed. It is only used to allow for varying the amount of sample while maintaining the correct flow rate; and, is used to prevent reagent from staining the inside of the cuvette.

## Slide 23
laser | detector |  | fluorochrome
405 |  |  | 
 | VL1 | 450/50 | Pacific Blue, Alexa Fluor 405, Brilliant Violet 421
 | VL2 | 522/31 | Horizon V500, LIVE/DEAD Aqua
 | VL3 | 603/48 | LIVE/DEAD Yellow, Qdot 605
488 |  |  | 
 | BL1 | 530/30 | FITC, GFP, Alexa Fluor 488, Calcein, LIVE/DEAD Green, ALDEFLUOR, HiLyte 488
 | BL2 | 574/26 | PE, propidium iodide, Hilyte 555
 | BL3 | 640 LP | PerCP, Pe-Cy7, PerCP-eFluor 710, LIVE/DEAD Red, 7-AAD
Attune (2 lasers, 6 detectors setup)

## Slide 24
Attune - Key performance features include:
Breakthrough acoustic technology that focuses cells or particles
Highest sample delivery rates commercially available (up to 1,000 μL/minute)
User-selectable collection rates
Equipped with: Blue/Violet: 488 nm (20 mW) and 405 nm (50 mW) lasers
8 parameters—6-color detection plus side scatter and forward scatter
User-changeable bandpass and dichroic filters
Simplified fluorescence compensation
Manual and automated compensation
Adjustable PMT voltage settings
Detection of up to 20,000 events/sec and 20 million events/file
Calibrated delivery volumes for volumetric analysis and absolute cell counts
Electronic resolution of 6 decades
Low fluid consumption (about 1 L/day); self-contained fluids
Countertop instrument—fits on standard lab bench or in laminar
Software:
No software licensing fees
Output file format: FCS 3.0
Live gating with automatic saving
User and administrator log-in

## Slide 25
Cancer
The Attune® Acoustic Focusing Cytometer, with its fast acquisition times and increased precision, overcomes the technological hurdles involved in analyzing CECs.
Stem Cells
Side Population Analysis
In this study, we demonstrate the ability of the Attune® Acoustic Focusing Cytometer with the blue/ violet laser configuration to quickly analyze a large number of events in search of very rare populations of stem cells.
Human Mesenchymal Stem Cells (hMSCs)
Adult human mesenchymal stem cells (hMSCs) are rare fibroblast-like cells capable of differentiating into a variety of cell tissues, including bone, cartilage, muscle, ligament, tendon, and adipose.
Cell Cycle Analysis
Cell cycle analysis is just one example of an application in which precise detection of differences in fluorescence intensity between multiple cell populations is critical...
Cell Proliferation
Successful proliferation analysis by dye dilution requires sensitive instrumentation and an extremely bright dye to accurately distinguish fluorescently labeled cells from autofluorescence after several cell divisions...
Marine Sample Analysis
Flow cytometry is a powerful tool for studying the biology, ecology, and biogeochemistry of marine photosynthetic picoplankton...
Immunophenotyping
The Attune® Acoustic Focusing Cytometer exhibits excellent segregation of populations in immunophenotyping experiments (with up to 6 colors)...
Apoptosis
The Attune® Acoustic Focusing Cytometer is compatible with a broad offering of reagents and kits for flow cytometric apoptosis testing...
GFP & RFP Detection
Data for GRP and RFP detection were collected from the Attune® Acoustic Focusing Cytometer using human osteosarcoma cells (U2OS) and BacMam CellLight® reagents...
Microbiological Applications
In recent years the application of flow cytometry to the study of various microbiological phenomena has increased, finding utility in studies that include detection and quantification...

## Slide 26
Example: Detecting human circulating endothelial cells
using the Attune® Acoustic Focusing Cytometer
Circulating endothelial cells (CECs) are mature cells shed from blood vessel
walls during the natural process of endothelial cell turnover.
Elevated levels of CECs have been reported in a host of pathological conditions including cardiovascular disorders, infectious diseases, immune disorders, post transplantation analysis, and cancer.
CECs are reported to be present in very low numbers: 0.01%–0.0001% of all peripheral blood mononuclear cells

## Slide 27
100
200
300
400
500
600
700
800
900
1000
Maximum Sample Input Rate (ml/min)
Instrument 1
Instrument 2
Instrument 6
Instrument 5
Instrument 4
Instrument 3
Attune®
Attune® Throughput Compared to Hydrodynamic Focused Instruments
Attune® can analyze at sample rates from 25µL/min to 1000µL/min without losing accuracy
Traditional Flow Cytometers can only run at most 150µL/min and will sacrifice data quality
Higher sample rates enable dilution of limited samples and analysis of Rare Events Faster
Hydrodynamic Focused Instruments
Notes: With Precious and Rare samples, the problem is that certain cells may not have enough volume of the original sample or cells may be fragile and can not withstand the centrifugation required to concentrate enough to run at low sample input rates. With the Attune® cytometer, being able to increase the sample rate up to 10x that of a traditional flow cytometer enables these customers to run rare or dilute samples. Also, because you can dilute your sample prior to running it on the Attune, customers with limited sample can save some of their sample for other analyses i.e. PCR, western, imaging etc.

## Slide 29
Attune souhrn
Výhody:
rychlost měření
jednoduché ovládání
sw licence bez omezení
snadná výměna emisních filtrů
nízká spotřeba nosné kapaliny (cca 1L denně)
Limitace:
jen dva lasery (6 barev)
pouze originální roztoky
dlouhý (i když automatický) shutdown
sw nedokáže importovat FCS data
nutnost nastavit určitý akviziční objem vzorku

## Slide 31
tlak nosné (oplašťující) kapaliny vede pufr kyvetou a vyšší tlak ve zkumavce se vzorkem zavádí vzorek do kyvety.
Princip hydrodynamického zaostření zarovná buňky v kyvetě „jako perly na šňůrce“ předtím než dojdou do bodu kde protnout paprsek laseru.
Hydrodynamické zaostření nemůže oddělit buněčné agregáty. Průtoková cytometrie vyžaduje suspenzi jednotlivých buněk!
Fluidika – shrnutí 2

## Slide 32: Image Stream & Flowsight Amnis – kombinace průtokové cytometrie a analýzy obrazu

## Slide 34: Amnis - aplikace

## Slide 35: Principy průtokové cytometrie a sortrování
sorting
zpracování signálu
analýza dat
kompenzace signálu
K. Souček Bi9393 Analytická cytometrie

## Slide 38
+++++
-----

## Slide 39
+++++
-----

## Slide 40
+++++
-----

## Slide 41
+++++
+++++
-----
+++++

## Slide 42
++
+++++
-----
+++++
++
+++++

## Slide 43
+
+++++
-----
+++++
+
+++++

## Slide 44
-----
+++++
-----
+++++
+++++

## Slide 45
-----
+++++
-----
+++++
+++++
-----

## Slide 46
- -
+++++
-----
- - - - -
- -
- - - - -

## Slide 47
-
+++++
-----
-
- - - - -
- - - - -

## Slide 48
-
+++++
-----
- - - - -
- - - - -

## Slide 49
-
+++++
-----
- - - - -
- - - - -

## Slide 50
-
+++++
-----
- - - - -
- - - - -

## Slide 51
http://www.cyto.purdue.edu/cdroms/cyto10a/seminalcontributions/fulwyler.html

## Slide 54
SORTING

## Slide 55
Frequency
Charge
Drop Delay
Amplitude
SORTING

## Slide 56
SORTING
Sheath Pressure
Nozzle Size
Frequency
Amplitude

## Slide 57
SORTING
Drop delay
Interrogation point
Breakoff

## Slide 58
SORTING
Each sort setup includes:
Sheath pressure
Breakoff window values
Side Stream window values
Instrument window > Laser tab values

## Slide 59
SORTING - Streams
Good
Bad

## Slide 60
SORTING – Setup Side Streams

## Slide 61: Drop Delay
interrogation point
drop delay
breakoff
Waste
BD FACS™ Accudrop technology
Accudrop beads
Diode laser
Camera
Optical filter
Notes: Before sorting, you need to make sure that you have an accurate drop delay setting. The Aria has integrated Accudrop technology which provides a simple way to adjust the drop delay.
The components used for Accudrop:
Diode laser: A low-power red laser which illuminates the lower portion of the stream when the sort block door is closed (there’s a safety interlock).
Camera: Provides an image of the side streams visible in this window. (located behind the round window in the sort block)
Optical filter: Used for viewing the fluorescence from the droplets containing Accudrop beads. (point out the button on this image that moves the filter into place)
Accudrop beads: Beads that are excited by the diode laser and emit within the range of the optical filter.
The drop delay is the distance between the laser intercept and the last connected drop, measured in time. The system needs to know when to charge the stream so that the intended droplet is charged before it is detached from the stream.
This diagram shows the diode laser and the lower camera window. When setting the drop delay, the waste aspirator drawer is blocking the collection tubes, so all beads will go to waste. Essentially, Accudrop allows us to view the accuracy of our sort in real-time, without having to do a post-sort analysis.

## Slide 62
Sorting - Sort Masks
Cells are randomized distributed over the stream

## Slide 63
Sorting - Sort Masks
Trailing
Interrogated
Leading

## Slide 64: Mask
A region of the stream monitored for the presence of cells
Determines how drops will be deflected if a sorting conflict occurs
Measured in 1/32 drop increments
Mask = 0
Mask = 8
Mask = 16
Mask = 32
4
4
8
8
16
16
Notes: At the laser intercept particles are defined as wanted or not based on the gates you created in the software. However, while the particle is traveling down the stream another decision is made on whether or not the particle will be sorted based upon the Sort Precision mode.
A sort precision mode is made up of a combination of masks. A mask is a region of the stream that will be considered to make a decision if there’s a sorting conflict. Each mask is measured in 1/32 drop increments.
The next few slides will discuss each individual mask, then we’ll look at how those three masks are combined into precision modes.
Note to instructor: Encourage customers to hold questions about putting the masks together until after you’ve discuss the basic definitions of all three.

## Slide 65: Conflict Resolution
Precision modes include three types of masks
  Yield
  Purity
  Phase
Notes: At the laser intercept particles are defined as wanted or not based on the gates you created in the software. However, while the particle is traveling down the stream another decision is made on whether or not the particle will be sorted based upon the Sort Precision mode.
A sort precision mode is made up of a combination of masks. A mask is a region of the stream that will be considered to make a decision if there’s a sorting conflict. Each mask is measured in 1/32 drop increments.
The next few slides will discuss each individual mask, then we’ll look at how those three masks are combined into precision modes.
Note to instructor: Encourage customers to hold questions about putting the masks together until after you’ve discuss the basic definitions of all three.

## Slide 66
Sorting - Sort Masks
Sort decisions are determined by sort masks
Target particles in a drop with 1/32-drop resolution

## Slide 67
Sorting - Yield Mask
The yield mask defines how many drops will be sorted Yield mask of 8/32 indicated in blue; target particle shown in green
Yield Mask

## Slide 68
Sorting - Purity Mask
Purity mask of 8/32 in blue, 4/32 in each adjacent drop; target particles in green, non-target particles in red
Purity Mask
Purity Mask

## Slide 71: Creation of a Voltage Pulse
Laser
Laser
Laser
Time
Voltage
Time
Voltage
Time
Voltage
1
2
3

## Slide 72: Height, Area, and Width
Time (µs)
Voltage
Pulse area(A)
Pulse Height (H)
Pulse Width (W)
0

## Slide 73: Signal processing
time
analog signal intesity
VOLTAGE
FSC ~ cell size
FL-1 (530/30nm) ~ green fluo.
FL-2 (585/42nm) ~ red fluo.
Analog/digital
conversion
Height
Width
Area ( ∫ )
FL- (H, W, A)
FL-1 (H)
FL-2 (H)
dot plot
0
1000
1000
Zesílení signálu (!)
lin nebo log
K. Souček Bi9393 Analytická cytometrie

## Slide 74
baseline
16, 364
Time
340
1985
7650
12420
Analog to Digital Converter

## Slide 75
Voltage In
PMT Power Supply
Levels 0–1000 Volts
Photon In
Signal
Out
Digital data to memory
Analog to Digital
Conversion
Digitize
the pulse
16,384 levels
Sample
the pulse
10 MHz
Analog to Digital Converter

## Slide 76
Parameters
Area: Sum of all height values
Height: Maximum digitized value X 16
Width: Area/Height X 64K
Data is displayed on 262,144 scale
282
3060
10270
358
4004
9568
14524

## Slide 80
24=16

## Slide 81: Lineární a logaritmický obvod
lineární
logaritmický
kompenzace fluorescenčního signálu

## Slide 82: AD převodníky
Počet bitů | # kanálů | rozlišení
8 | 256 | 39.1 mV
10 | 1024 | 9.77 mV
12 | 4096 | 2.44 mV
14 | 16384 | 610 mV
16 | 65536 | 153 mV
18 | 262144 | 38.1 mV
20 | 1048576 | 9.54 mV
22 | 4194304 | 2.38 mV
24 | 16777216 | 596 nV
28 = 256
210 = 1024
.
.
.
Full scale measurement range = 0 to 10 volts
ADC resolution is 12 bits: 212 = 4096 quantization levels
ADC voltage resolution is: (10-0)/4096 = 0.00244 volts = 2.44 mV
K. Souček Bi9393 Analytická cytometrie

## Slide 83: Kolik bitů?
Pokud konvertujeme analogový signál pomocí 8 bitového ADC – máme 256 kanálů (28=256) odpovídajících rozsahu 0-10 V
Rozdíl mezi kanály je 10/256=~40mV
0 50 100 150 200 250
10V
1V
100mV
Channels
upraveno podle J.P.Robinson

## Slide 84: Ideální logaritmický zesilovač
0 50 100 150 200 250
10 V
1 V
100 mV
0 50 100 150 200 250
10 V
1 mV
Channels
Linearní
Log
1 V
100 mV
10 mV
Log amp
upraveno podle J.P.Robinson

## Slide 85: Logaritmické zesílení & dynamický rozsah
upraveno podle J.P.Robinson
lin
log

## Slide 86: Kompenzace fluorescenčního signálu
…později

## Slide 87: Analýza dat
Zobrazení dat
  histogram
  dot plot
  isometric display
  contour plot
  chromatic (color) plots
  3 D projection
Gating

## Slide 88: Způsoby pro zobrazení dat
K. Souček Bi9393 Analytická cytometrie

## Slide 89: Histogram distribuce četnosti
Histogram zobrazuje četnost částic pro jeden parametr
Jednoduchý výstup
Nekorelujeme s dalším parametrem
Problém s identifikací populací
K. Souček Bi9393 Analytická cytometrie

## Slide 90: Dot plot
Zobrazuje korelaci dvou libovolných parametrů
Jednotlivé tečky představují konkrétní změřené buňky (částice)
Hodnoty pro řadu částic mohou ležet ve stejném místě
Nemáme informaci o relativní denzitě částic
Problémy s vykreslením v případě velkých objemů dat
K. Souček Bi9393 Analytická cytometrie

## Slide 91: Density & contour plot
Density plot:
Zobrazuje dva parametry jako frekvenci četnosti
barva a nebo její odstín odpovídá četnosti částic
Contour plot:
spojnice spojuje body (částice) se stejnou hodnotou signálu
V podstatě simulujeme 3D graf – třetí rozměr je frekvence
K. Souček Bi9393 Analytická cytometrie

## Slide 92: Čas jako jeden z parametrů
K. Souček Bi9393 Analytická cytometrie

## Slide 93: 3D zobrazení
2 parametry + četnost
3 parametry společně
K. Souček Bi9393 Analytická cytometrie

## Slide 94: 3 Color Combinations
Negatives
Positives
4+4=8
FITC
PE
APC
4+4+4=12
upraveno podle J.P.Robinson

## Slide 95: 3 Color Combinations
FITC
PE
APC
4+4+4=12
PE-FITC Pattern
FITC
TR-FITC Pattern
FITC
TR-PE Pattern
Phycoerytherin
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
.1
1
10
100
1000
CD4
CD5
CD38
K/L
CD45
CD3
Negative
Negative
CD10
HLA-DR
CD20
CD8
CD7
CD2
CD8
CD8 (dim)
CD2
Negative
CD4
CD5
K
Texas Red
Texas Red
Phycoerytherin
CD3
CD3
IgM
IgG
CD3
CD8
CD3
CD20
IgG
L
CD20
CD20
CD45
upraveno podle J.P.Robinson

## Slide 96: „Gating“
Real-time gating vs. softwarový „gating“
Určení regionů
Strategie „gatingu“
Analýza kvadrantů
Boolean „gating“
zpětný „gating“

## Slide 97: Real-Time vs. Software Gating
Real-time (live) gating:
-omezuje akceptovaná data během měření
Software (analysis) gating:
-vyřazuje určitá data během následné analýzy
upraveno podle J.P.Robinson

## Slide 98: Určení regionů
Objektivní nebo subjektivní?
  -školení/schopnosti/trénink
Možné tvary:
  -obdelník
  -elipsa
  -“free-hand“ (polygon)
  -kvadrant
Statistika
- počet
- podíl (%)
- průměr, medián, S.D., CV, ….

## Slide 99: Region vs. gate
Region
oblast (plocha) v grafu definovaná uživatelem
mnoho regionů v jednou grafu
ohraničujeme pomocí nich populace našeho zájmu
je možné je barevně odlišit
je definován stejně pro všechny vzorky v analýze
Gate
je definován jako jeden a nebo více regionů zkombinovaných pomocí logických operátorů (AND, OR, NOT; Booleova logika)
K. Souček Bi9393 Analytická cytometrie

## Slide 100
Region 1 established
Gated on Region 1
Using Gates
log PE
upraveno podle J.P.Robinson

## Slide 101: Statistika
Aritmerický průměr
Geometrický průměr
Medián
– odhad střední hodnoty
– není ovlivněn extrémními hodnotami
Směrodatná odchylka
Koeficient variance
Modus – nejčastější hodnota
K. Souček Bi9393 Analytická cytometrie

## Slide 102: Statistika pro histogram
K. Souček Bi9393 Analytická cytometrie

## Slide 103: Analýza kvadrantů
(+ +)
( - +)
(+ -)
(- -)
K. Souček Bi9393 Analytická cytometrie

## Slide 104: Logický „Gating“ (Booleova logika)
S překrývajícími se oblastmi máme mnoho možností:
upraveno podle J.P.Robinson

## Slide 105: Boolean Gating
Not Region 1:
upraveno podle J.P.Robinson

## Slide 106: Boolean Gating
Not Region 2:
upraveno podle J.P.Robinson

## Slide 107: Boolean Gating
Region 1 or Region 2:
upraveno podle J.P.Robinson

## Slide 108: Boolean Gating
Region 1 and Region 2:
upraveno podle J.P.Robinson

## Slide 109
Not (Region1 and Region 2):
Boolean Gating
upraveno podle J.P.Robinson

## Slide 110
90 Degree Scatter
0
200
400
600
800
1000
8
15
20
30
40
50
100
200
1000
Lymphocytes
Monocytes
Neutrophils
Side Scatter Projection
Light Scatter Gating
Scale
upraveno podle J.P.Robinson

## Slide 111: Back Gating
Region 4 established
Back-gating using Region 4
log PE
Back gate
upraveno podle J.P.Robinson

## Slide 112: Back Gating
K. Souček Bi9393 Analytická cytometrie

## Slide 113
FITC+
PE+
FITC+
APC+
PE +
APC+
3 Parameter Data Display
upraveno podle J.P.Robinson

## Slide 114: Vícebarevné analýzy generují mnoho dat…
1 2 3 4 5 6 7 8 9 10
3 color
4 color
5 color
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
Log Fluorescence
QUADSTATS
Log Fluorescence
++
--
+-
-+
upraveno podle J.P.Robinson

## Slide 116: Detection and monitoring of normal and leukemic cell populations with hierarchical clustering of flow cytometry data
Cytometry Part A Volume 81A, Issue 1, pages 25-34, 11 OCT 2011 DOI: 10.1002/cyto.a.21148 http://onlinelibrary.wiley.com/doi/10.1002/cyto.a.21148/full#fig1
Notes: HCA‐Main hematopoietic populations in normal BM. A: Dendrogram with heatmap‐HCA of 104 ungated events acquired from normal BM. Heatmap shows relative levels of all eight parameters (columns) in all 104 events (rows) in color coding (blue, low expression; red, high expression). Dendrogram shows the hierarchy of cells based on their similarity in all parameters measured. The x‐axis under the dendrogram represents similarity distance. Colored branches of the dendrogram are selected clusters, as displayed in B. B: The main populations (as identified by HCA) are displayed on a conventional forward versus side scatter dot plot. C: All two‐parameter combination plots of a defined cluster (marked by red rectangle in A). This population is negative for most of the antibodies from the B cell panel and may be difficult to detect by standard gating as it overlaps with other populations in all 28 two‐parameter plots. Nevertheless it is a compact, homogenous population most likely of T cell origin. This red population was drawn on top of the remaining cells (gray). D: Mirror image to C. The gray population was drawn on top of the red population. Fluorescent dyes used were: CD10‐FITC, CD22‐PE, CD117‐PerCP‐Cy5.5, CD38‐PE‐Cy7, CD34‐APC, CD19‐APC‐Cy7, and in all cases pulse height was used.
© This slide is made available for non-commercial use only. Please note that permission may be required for re-use of images in which the copyright is owned by a third party.

## Slide 118: The Flow Cytometry: Critical Assessment of Population Identification Methods (FlowCAP)
The goal of FlowCAP is to advance the development of computational methods for the identification of cell populations of interest in flow cytometry data. FlowCAP will provide the means to objectively test these methods, first by comparison to manual analysis by experts using common datasets, and second by prediction of a clinical/biological outcome.

## Slide 119: Způsoby pomocí kterých lze upravit výsledky:
1. Odstranění „doublets“
2. Čas jako parametr pro kontrolu kvality
Příklad - pro DNA analýzu je třeba:
- odstranit „debris“ a shluky
- odstranit „doublets“
- udržovat konstantní průtok
K. Souček Bi9393 Analytická cytometrie

## Slide 120
DNA Histogram
G0-G1
S
G2-M
Fluorescence Intensity
# of Events

## Slide 121: Co je problém při vícebarevné detekci?
K. Souček Bi9393 Analytická cytometrie

## Slide 122
Emission Spectra–Spectral Overlap

## Slide 123: Kompenzace fluorescenčního signálu při vícebarevné detekci
Proces při kterém dochází k eliminaci všech fluorescenčních signálů kromě signálu z fluorochromu který má být na příslušném detektoru detekován
Nastavení pomocí mixu mikročástic či buněk označených/neoznačených příslušnými fluorochromy.
K. Souček Bi9393 Analytická cytometrie

## Slide 124: Co je problém při vícebarevné detekci?
K. Souček Bi9393 Analytická cytometrie

## Slide 125: Kompenzace fluorescenčního signálu při vícebarevné detekci
K. Souček Bi9393 Analytická cytometrie

## Slide 126
“Bright” = good resolution sensitivity

## Slide 127
Various fluorochromes-stain index

## Slide 128
Choices for 6,- 8,- 10,- and more colors

## Slide 129
Fluorochrome selection considerations

## Slide 130
FITC Spillover
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42
750nm
800nm

## Slide 131
FITC Compensation
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42

## Slide 132
FITC Compensation
650nm
700nm
PerCP-Cy5.5
695/40
500nm
600nm
FITC
530/30
Relative Intensity
Wavelength (nm)
550nm
PE
585/42

## Slide 133
FITC Compensation
Dot plot showing uncompensated FITC data
Dot plot showing compensated FITC data
Biexponential dot plot showing compensated FITC data

## Slide 134
FITC Spillover
600nm
500nm
550nm
FITC
530/30
Relative Intensity
Wavelength (nm)
PE
585/42

## Slide 138: Kompenzace fluorescenčního signálu
K. Souček Bi9393 Analytická cytometrie
FITC positive & negative
PE negative beads
#2

## Slide 139: Kompenzace fluorescenčního signálu
K. Souček Bi9393 Analytická cytometrie
FITC positive & negative
PE negative beads
NONE!

## Slide 140: Kompenzace fluorescenčního signálu

## Slide 141: Nastavení kompenzací
parametr - detektor amp.
FL1 - 544
FL2 - 434
kompenzace
FL1 - 1.1%FL2
FL2 - 17.5%FL1
značené mikročástice – pro běžně konjugované fluorochromy
značené buňky – pro vitální značení

## Slide 142
Effects of Changing PMT Values
FITC Voltage Increased by 5 V
FITC Voltage Decreased by 5 V
Correct Compensation

## Slide 143
Which marker for compensation?
Small errors in compensation of a dim control (A) can result in large compensation errors with bright reagents (B & C).
Use bright markers to setup proper compensation.

## Slide 144
BD Comp Beads
Always positive
Bright staining
Save sample (HIV patients)
Use the same antibody for compensation and the real experiment

## Slide 145
BD Comp Beads

## Slide 146
PBMC were stained as shown in a 3-color experiment. Compensation was properly set for all spillovers
Courtesy Mario Roederer
Fluorescence Minus One

## Slide 147: Tandemové značky

## Slide 148: Tandemové značky - příklad

## Slide 149
Tandems are light sensitive
0 hours
2 hours
22.5 hours
PE
(FL2)
CD8
CD3
PE-Cy5
PE-Cy7
Time Sample Left in Light

## Slide 150: Kompenzace - literatura
Mario Roederer - Compensation in Flow Cytometry
Current Protocols in Cytometry (2002) 1.14.1-1.14.20 John Wiley & Sons, Inc.
M. Loken, D. R. Parks, & L. A. Herzenberg (1977). Two-color immunofluorescence using a fluorescence-activated cell sorter. J. Histochem. Cytochem. 25:899-907. M. Roederer & R. F. Murphy (1986). Cell-by-cell autofluorescence correction for low signal-to-noise systems: application to EGF endocytosis by 3T3 fibroblasts. Cytometry 7:558-565. S. Alberti, D. R. Parks, & L. A. Herzenberg (1987). A single laser method for subtraction of cell autofluorescence in flow cytometry. Cytometry 8:114-119. C. B. Bagwell & E. G. Adams (1993). Fluorescence spectral overlap compensation for any number of flow cytometry parameters. in: Annals of the New York Academy of Sciences, 677:167-184.
K. Souček Bi9393 Analytická cytometrie

## Slide 151
No Data Analysis Technique Can Make Good Data Out of Bad Data!
Shapiro’s 7th Law of Flow Cytometry

## Slide 152: Shrnutí přednášky
sorting
zpracování signálu
vizualizace dat a „gating“
kompenzace
Na konci dnešní přednášky byste měli:
Znát základní principu sortování,
popsat způsob zpracování signálu,
rozumět lin / log zesílení signálu,
rozeznat jednotlivé způsoby vizualizace dat,
chápat základní principy „gatingu“,
znát princip kompenzace signálu při vícebarevné detekci.
K. Souček Bi9393 Analytická cytometrie