## Slide 1
S5015
Light microscopy methods in biology
Lecture 1: History and principles of light microscopy
Milan Ešner
Cellular Imaging Core Facility, Ceitec MU
milan.esner@ceitec.muni.cz

## Slide 2: S5015 : Light microscopy methods in biology
Lecture 1: September 30, 2021 at 9:00
History of microscopy, fundamentals of light microscopy. Contrasting techniques in transmitted light microscopy – brightfield, darkfield, phase contrast, DIC.
Widefield fluorescence microscopy - deconvolution, point spread function, point spread function measurements, description.
Detectors in widefield microscopy, filters, light sources.
Lecture 2: October 14, 2021 at 9:00
Confocal microscopy - principles, types of confocal microscopes – scanning, spinning disc, channel vs spectral imaging, linear unmixing.
Lasers, detectors, beam splitters, optical sections. Lightsheet microscopy.
Lecture 3: October 25, 2021 at 9:00
Imaging below diffraction limits – SIM, SMLM, STED
Lecture 4: November 4, 2021 at 9:00
Basics of Image analysis – image types, pixels, voxels, basic filters, image processing, metadata, deconvolution, 3D visualization,
objects detection, quantification.
Practical 1: December 14, 2021 at 9:00
Transmitted light microscopy, widefield fluorescence, confocal, Apotome, SIM
Practical 2: December 15, 2021 at 9:00
Sample preparation, single molecule localization microscopy acquisition and analysis, deconvolution.

## Slide 3: Antonie van Leeuwenhoek (1632 – 1723)
Dutch businessman and scientist. Best known for his pioneering work in microscopy. Used single lens microscope
of his own design.

## Slide 4: Robert Hooke (1635 – 1703)
English scientist, philosopher, architect and polymath. Improved design of compound microscope and use it
for scientific exploration – book Micrographia. Use term <cell> in biological context.

## Slide 5: Carl Zeiss (1816 – 1888)
German scientific instrument maker, optician and businessman. Founder of Carl Zeiss workshop in 1864
together with Ernst Abbe and Otto Schott in Jena, Germany.

## Slide 6: Ernst Abbe (1840 – 1905)
German physicist, optical scientist, entrepreneur and social reformer. With Zeiss and
Schott he developed numerous optical instruments. Among his inventions are apochromatic lens,
Abbe condenser, refractometer.
He defined numerical aperture -> mathematically describing optical resolution of light microscopes.
He was also strongly implicated in unions and workers rights. In Zeiss he introduced 8-hour workday.

## Slide 7: George Biddell Airy (1801 – 1892)
English mathematician and astronomer. He established Greenwich as the location
of the prime meridian. His main contribution to microscopy was explaining of
formation of concentric diffraction pattern around the bright spot of light,
named after him – Airy disc.

## Slide 8: August Koehler (1866 – 1948)
German optical scientist. He developed and improved light illumination of sample, Koehler illumination.

## Slide 9: Siedentopf and Zsigmondy
First concept of light sheet microscope – illumination objective orthogonally to detection objective. Zsigmondy received Nobel prize for chemistry in 1925 for his research on colloid solutions.

## Slide 10: Frits Zernike (1888 – 1966)
Dutch physicist. Invented phase contrast microscopy. Producing high contrast images of
transparent samples, hardly visible in standard bright field transmitted light imaging.
Widely used technique of contrast enhancing. Received Nobel prize in 1953.
Georges Nomarski
(1919-1997)
Polish physicist. Developed differential interference contrast (DIC) –
contrast enhancing technique based on polarized light.

## Slide 11: Marvin Minsky (1927-2016)
American scientist, doing mainly research of artificial intelligence. He is
Co-founder of Massachusetts Institute of Technology (MIT).
He developed first scanning confocal microscope in 1957.

## Slide 12: Mojmír Petráň (*1923)
Czechoslovak scientist. Studied medicine at Charles university, Prague.
With Milan Hadravsky he developed spinning disc confocal microscope in 1964, using Nipkow disc from TV.

## Slide 13: Jan Huisken, Eric Stelzer
Developed fluorescence light sheet microscope (selective plane illumination microscope) in 2004.
Almost 100 years after Siedentopf and Zsigmony microscope they built first fluorescence microscope,
using different optical paths for illumination and detection. Today exists many variants of lightsheet
microscopes and several commercial systems are available.

## Slide 14: Mats Gustafsson (1960-2011)
Developped method of structured illumination microscopy SIM. SIM surpass diffraction limit of optical microscopes
and is able to achieve resolution below 100 nm.

## Slide 15: Eric Betzig (*1960)
Inventor of single molecule localization technique STORM/PALM together with William Moerner and
lattice light sheet technology. Received Nobel prize in 2014 together with Eric Betzig and Stephen Hell

## Slide 16: William Moerner
Contributed to development of sinlge molecule localization microscopy. Received Nobel prize in 2014 together
with Eric Betzig and Stephen Hell

## Slide 17: Steffen Hell
Invented stimulated emission depletion (STED) microscopy. Received Nobel prize in 2014 together
with Eric Betzig and Stephen Hell

## Slide 18
Principles of light microscopy

## Slide 19
In microscopy you don’t observe an OBJECT, you observe an IMAGE of the object
1) Light diffraction by the specimen
2) collection of diffracted/non-diffracted rays by objective
3) interference of diffracted and non-diffracted rays in the image plane.
Three steps in image formation
Magnification
Resolution
Contrast
Image formation

## Slide 20
Monochromatic - single wavelength (color).
Polychromatic
Coherent – waves of the same wavelength that maintain the same phase relation.
Noncoherent
Polarized – waves with E-vectors in parallel.
Nonpolarized
Collimated – waves with the same coaxial paths of propagation (same direction).
Divergent
Visible electromagnetic radiation with electric and magnetic fields whose amplitudes
oscillate as a sine function over dimensions of space and time.
Dual character: waves – amplitude and phase
particles – photons.
shorter wavelength = higher energy
Light

## Slide 21
Light diffraction
When light passes at the edge of object (opening, aperture, specimen),
light is bended on edges and passes to the area of geometrical shadow.
Due to the diffraction, Image of a point is not a point

## Slide 22
Constructive interference
Light diffraction and interference

## Slide 23
Destructive interference
Light diffraction and interference

## Slide 24
0
1st
1st
2nd
2nd
3nd
3nd
Light diffraction and interference

## Slide 25
Air objective
Oil immersion
objective
0.95
1.515 * 0.92
1.4
Immersion objectives have higher NA and therefore higher resolution.
Adjust refractive index of immersion to your medium!!!
Mowiol, prolong, vectashield –> oil immersion
Glycerol -> glycerol immersion
PBS, medium –> water immersion
Mismatch = high spherical aberrations
From P. Evennett
www.zeiss-campus.magnet.fsu.edu
Numerical aperture

## Slide 26: Refractive index
Medium | R. Index (n)
Air | 1
Water | 1.333
Glycerol | 1.46
Immersion Oil | 1.515
Cover glass | 1.515
- Determines how much light is bent (refracted), when entering the material
n1 * sin(1) = n2 * sin(2)
Snell’s law:
n1
n2
1
2
RI is different for every wavelength – light dispersion on prism.
air
water
Light entering to different medium changes the speed
Refraction

## Slide 27: Ernst Karl Abbe
(1840 – 1905)
d =
Microscope
Numerical aperture
Resolution

## Slide 28
By Spencer Bliven - Own work, Public Domain, https://commons.wikimedia.org/w/index.php?curid=31456019
Two point sources are regarded as just resolved when the principal diffraction maximum of one image coincides with the first minimum of the other.
Rayleigh criterion
R
R: resolution/distance
λ: wavelength of light
NA: numerical aperture
Resolution

## Slide 29
Magnification does not play a role in resolution!
Only NA and  are taken into account!
For green light (550 nm and 1.4 NA objective) d  240 nm
Limiting resolution of light microscopes (diffraction limited).
Objective | Resolved distance d (nm)
5x/0.15 | 2200
10x/0.3 | 1100
20x/0.5 | 700
40x/0.75 | 450
40x/1.3 | 260
63x/1.4 | 240
100x/1.3 | 260
Confocal microscope
1 AU pinhole
Confocal microscope
0.25 AU pinhole
Conventional microscope
Lateral resolution

## Slide 30
⍺
Objective coding

## Slide 31
Chromatic aberrations
Spherical aberrations
F
F
F
F
Most frequent lens abberations
Apochromatic objectives
Use correct coverslip
Refraction media match

## Slide 32
Transmitted light microscopy

## Slide 33
Conjugated planes
Field set
Aperture set

## Slide 34
Light passing directly through specimen to eys ot other detector
BF – bright field, some microscopes H on condensor
+ Easy to setup, no special requirements
- Low contrast
Bright field
Closed illuminating aperture

## Slide 35
Only diffracted light from specimen passing to objective
DF – bright field
+ Easy to setup, no special requirements, for high NA
Special condenser required
Dark field
Dark field

## Slide 36
Transparent objects not visible in the transmitted light mode still
diffract light path and cause phase shift
Phase contrast imaging transform phase differences into
amplitude differences.
Ph1, Ph2, Ph3 – phase contrast
+ Easy to setup, nice contrast, working on plastic dish
- Special objectives required, halo effect around objects
Phase contrast

## Slide 37
Differential interference contrast (DIC)
Based on polarised light. Transform local gradient in optical path
length (refractive index and thickness) in an object into regions of
contrast in the object image. Amplitude differences in the image
due to refractive index and/or thickness of the specimen.
+ nice contrast, 3D perspectives, no halos around objects
- Works on glass only, not trivial to setup
https://www.olympus-lifescience.com

## Slide 38
Reflected light
Fluorescence microscopy
Emisní
filtr
Excitation
filter
light

## Slide 39
Fluorescein
Fluorescence

## Slide 40
widefield (epifluorescence)
confocal microscopes
  point scanning microscopes
  spinning disc microscopes
lightsheet microscopes
Diffraction limited resolution
Resolution below diffraction limit - nanoscopes
Structured illumination (SIM)
Stimulated emission depletion (STED)
Stochastic optical reconstruction (STORM)
Types of fluorescent microscopes

## Slide 41
Fluorescein
https://en.wikipedia.org/wiki
Difference between positions of the band maxima of the absorption and emission spectra
Stokes shift
Fluorescence

## Slide 42
Fluorescein
Jablonski diagram
S0
S1
S2
ground state
excited state
energy
Radiation - fluorescence
Fluorescence

## Slide 43
George Biddell Airy (1801-1892)
Best focused point source of light won’t appear as a point, but as a spot
surrounded by series of concentric bright rings – Airy pattern.
Size of Airy disc depend on the light wavelength and the size of aperture.
Imaging
Device
(circular aperture)
Airy disc (d)
- Airy disc contain ~84% of the light from the point source
Airy disc

## Slide 44
Three dimensional diffraction pattern of light emitted from an infinitely small point source
in the specimen and transmitted to image plane
object
Point Spread Function (PSF)
Image
convolution
Point spread function

## Slide 45
Three dimensional diffraction pattern of light emitted from an infinitely small point source
in the specimen and transmitted to image plane
object
Point Spread Function (PSF)
Image
deconvolution
Deconvolution
convolution

## Slide 46
https://en.wikipedia.org/wiki/Fluorescence
Fluorescent proteins in living organisms
Roger
Tsien
Nobel prize in chemistry 2008
Discovery and development
of the green fluorescent protein
GFP
Martin
Chalfie
Osamu
Shinomura
Aequorea victoria
Fluorescence

## Slide 47
Excitation light
Emission spectrum
Sample
Excitation
filter
Emission
filter
Dichromatic
mirror
Detector
Components of fluorescent microscope

## Slide 48
Single bandpass (limitations min and max)
Longpass filter
Shortpass filter
Multi bandpass filter
Filters and beam splitters

## Slide 49
Crosstalk/bleed through
Filters and beam splitters

## Slide 50
Mercury short arc lamp:
Obsolete, short life time – 200 hours, toxic wapors
Manual alignmnent
Light sources

## Slide 51
Metal halide lamp:
Longer life time – 2000 hours
No alignment, contains also mercury
Light sources

## Slide 52
Xenon lamp
Light sources

## Slide 53
LED – light emmiting diodes
Single wavelength, or combination of several LED‘s
Instantly switching ON/OFF
Light sources

## Slide 54
Widefield microscopes – CCD or sCMOS sensors
Pixel number, pixel size, bit depth, quantum efficiency, spectral response
photons
electrons
photodiode
Photoactive part
(photodiode)
Isolated interline device
(charge transfer)
- Smallest unit of sensor with defined size (x,y dimensions)
- Sensor consists of many pixels
Detectors

## Slide 55
Binning and resolution
Several neighborhoods pixels are accumulated together and readed as one unit
Increasing sensitivity, decreasing resolution, decreasing size of image
Binning 1 – full resolution
Resolution = 12 x 12 pixels
Pixel size = 6 x 6 μm
Binning 2
Resolution = 6 x 6 pixels
Pixel size = 12 x 12 μm
Binning 4
Resolution = 3 x 3 pixels
Pixel size = 24 x 24 μm
Pixels
Detectors

## Slide 56
Spectral response and quantum efficiency
Detectors

## Slide 57
Analog to digital converter (ADC) – convert analog signal (current in Voltage) to discrete
digital values - number of grey values.
Number of bits | 1 | 2 | 4 | 8 | 12 | 14 | 16
Number of grey values | 2 | 4 | 16 | 256 | 4,096 | 16,384 | 65,536
Bit depth
Detectors

## Slide 58
Correct exposure – read histogram
Exposure

## Slide 59
1 sec
2 sec
4 sec
8 sec
Saturation - blooming
Correct exposure – read histogram
Exposure

## Slide 60
Important to know resolution of the system, avoid oversampling and undersampling
Pixel size should be 2-3 x smaller than the actual resolution of the system
Objectiv: Zeiss Plan-Apochromat 100x/1.4
Optovar:1.6x
C-mount: 1x
Camera: Hamamatsu Flash 4.0
Resolution (GFP) = 0.51 * 550/1.4 = 200 nm
Camera pixel size = 6.5 x 6.5 μm
Magnification = 100 x 1.6 x 1 = 160
Image pixel size = 0.0406 μm = 40.6 nm
Oversampled: Max. resolution = 200 nm, image pixel = 40 nm
Nyquist requirement = 2.3 times the resolution = 85 nm
Digital sampling

## Slide 61
Important! Acts as external part of the objective
Standard thickness is 0.17 mm (170 m) == grade 1.5, best 1.5H (high precision 170 m ± 0.05 m)
Variability should not exceed 0.1 for objective with high NA
Some objectives have correction collar to adjust for different thickness of cover slip
Grade Number | Thickness (mm)
0 | 83-130
1 | 130-160
1.5 | 160-190
2 | 190-250
Refraction index should be close to refraction index of glass == 1.5168
Home made – Mowiol solution
Commercial solutions – VectaShield, Prolong glass
Mounting media
Coverslips

## Slide 62
Thank you for your attention
Lecture 2: October 14, 2021 at 9:00
Confocal microscopy - principles, types of confocal microscopes – scanning, spinning disc, channel vs spectral imaging, linear unmixing.
Lasers, detectors, beam splitters, optical sections. Lightsheet microscopy.